## Results
CDMAM Analysis Results
Details |  |  | Detailed threshold gold thickness results
CDMAM version: |  | 3.4 | Disc diameter* (mm) | Automatic Threshold Thickness (µm) | Predicted Human Gold Thickness (µm) | Fitted Threshold Gold Thickness (µm) | SE (Fitted Threshold (µm) | Limits in protocol
CDCOM version: |  | v1.6.0 | [5 empty] | Acceptable level (µm) | Achievable level (µm)
Date: |  | 2019-10-21
Centre: |  | …..
|  |  | 0.1 |  |  |  |  | 1.68 | 1.1
Manufacturer: |  | ….. | 0.13
Model: |  | ….. | 0.16
Number of images: |  | 8 | 0.2
|  |  | 0.25 |  |  |  |  | 0.352 | 0.244
Image type: |  | ….. | 0.31
Target / Filter: |  | ….. | 0.4
kV: |  | ….. | 0.5 |  |  |  |  | 0.15 | 0.103
0.63
0.8
|  |  | 1 |  |  |  |  | 0.091 | 0.056
The error bars on this graph show 95% (2 times SE) confidence intervals for predicted data.