## Results
CDMAM Analysis Results
Details |  |  | Detailed threshold gold thickness results
CDMAM version: |  | 4.0 | Disc diameter (mm) | Automatic Threshold Thickness (µm) | Predicted Human Gold Thickness (µm) | Fitted Threshold Gold Thickness (µm) | SE (Fitted Threshold (µm) | Limits in protocol
CDCOM version: |  | v1.0.0 | [5 empty] | Acceptable level (µm) | Achievable level (µm)
Date: |  | 2019-10-21
Centre: |  | …..
|  |  | 0.08 | 1.2687 | 1.8391 | 1.235
Manufacturer: |  | ….. | 0.09 |  |  | 0.913
Model: |  | ….. | 0.1 | 0.74801 | 1.1502 | 0.756 | 0.13699 | 1.68 | 1.1
Number of images: |  | 8 | 0.13 | 0.41989 | 0.68292 | 0.534 | 0.073947
|  |  | 0.15 |  |  | 0.447
Image type: |  | ….. | 0.18 | 0.26631 | 0.45265 | 0.353 | 0.052704
Target / Filter: |  | ….. | 0.21 | 0.1948 | 0.3441 | 0.286 | 0.037909
kV: |  | ….. | 0.25 | 0.13414 | 0.24585 | 0.225 | 0.027942 | 0.352 | 0.244
|  |  | 0.3 | 0.096808 | 0.18299 | 0.175 | 0.021615
|  |  | 0.35 |  |  | 0.143
|  |  | 0.42 | 0.077603 | 0.15125 | 0.132 | 0.017322
|  |  | 0.5 | 0.055947 | 0.11116 | 0.121 | 0.015308 | 0.15 | 0.103
|  |  | 0.57 |  |  | 0.1
|  |  | 0.66 | 0.042914 | 0.086493 | 0.087876 | 0.013972
|  |  | 0.77 |  |  | 0.08
|  |  | 0.88 | 0.033787 | 0.069759 | 0.069893 | 0.01328
|  |  | 1 | 0.026926 | 0.056922 | 0.056858 | 0.012907 | 0.091 | 0.056
The error bars on this graph show 95% (2 times SE) confidence intervals for predicted data.